WI-NETRA  ·  HEALTH DOMAIN EXTENSION  ·  v2

Breathing-Rate Monitoring & Respiratory Anomaly Detection

Product-shaped module specification · feasibility-scoped for a 6-week internship

Build priority tags:  CORE = must work for the demo   ·   SUPPORTING = product shell, reuses FYP patterns   ·   STRETCH = prototyped only if time permits

# Overview & Architecture

This is the product-shaped specification for the health-domain extension of Wi-Netra: a WiFi CSI–based product that estimates a person's breathing rate in real time and flags respiratory anomalies, running on dedicated ESP32-S3 hardware and a separate Firebase project, fully decoupled from the Wi-Netra intrusion-detection system.

Breathing rate is treated as a measurement problem, not a classification problem: respiration is a periodic signal in a known band (0.1–0.5 Hz), so the rate is measured directly with signal processing rather than learned by a model. Machine learning is reserved for optional anomaly detection, where it can earn its place. Compute is split across three layers so each stage runs where it is cheapest to build and most robust:

- On-device (ESP32-S3): CSI capture, the DSP breathing-rate estimator (with a confidence/validity output), and deterministic rule-based anomaly flags. Only derived values leave the device — raw CSI is processed and discarded locally, preserving privacy.

- Cloud (Firebase): authentication and a Realtime Database that relays each device's derived breathing rate and status to subscribers in near real time. No local MQTT broker and no Raspberry Pi.

- Apps (phone + web): a single viewer app with a multi-patient switcher, an optional on-phone anomaly model, push notifications, and a small web admin panel for oversight.

Product model: one App User account (the same experience whether the user is the patient or a caregiver) can be linked to several devices — each device paired to one patient — and switches between them like a real-time multi-patient monitor. A separate Admin role manages users, devices, and device health.

| ## 1. User Management & Authentication | SUPPORTING |
| --- | --- |

Adapted from Wi-Netra 6.1 — slimmed to two roles, on a separate Firebase project.

Description

Manages accounts with just two roles, separated by function rather than title. Patients and caregivers share one identical App User experience — the only difference is which device(s) an account is linked to — so no separate “patient” vs “caregiver” role is needed. A distinct Admin role operates the fleet.

Key Features

- Two roles only: App User (patient or caregiver — identical capabilities) and Admin.

- An App User account can be linked to multiple devices, each device paired to one patient, enabling the in-app patient switcher.

- Registration, login, logout, password reset, email verification.

- Firebase Authentication with ID tokens; access scoped to the devices linked to the account.

- Runs on an isolated Firebase project so health and intrusion data never mix.

| ## 2. CSI Acquisition & Device Setup (ESP32-S3) | CORE |
| --- | --- |

Health counterpart of Wi-Netra 6.2 — dedicated hardware, with product onboarding folded in.

Description

Captures raw Channel State Information from a dedicated ESP32-S3 transmitter/receiver pair placed near the subject. The ESP32-S3 is chosen over the ESP32-C6 because its CSI driver reports the full LLTF + HT-LTF stack correctly, whereas the C6 driver currently omits L-LTF data and returns subcarriers in the wrong order — a documented Espressif firmware bug. Correct, dense subcarriers are exactly what breathing-rate signal quality depends on, and the S3's dual-core SIMD gives DSP headroom. A short guided setup step makes deployment feel like a product, not a lab rig.

Key Features

- Dedicated ESP32-S3 TX/RX pair; full LLTF + HT-LTF CSI capture for subcarrier diversity.

- 10 Hz sampling (well above the ~2 Hz Nyquist need for respiration) with 30 s windows and 15 s stride — the rate used by published ESP32 respiration systems, and comfortable on the S3's compute/RAM budget.

- Fixed geometry for best signal: TX/RX ~1–2 m apart, subject seated between them, chest broadly perpendicular to the line-of-sight path.

- Steady packet rate via a fixed-rate transmitter (or router-ping), with linear interpolation to correct residual packet-rate jitter — a known ESP32 CSI issue with a standard fix.

- Guided onboarding / baseline calibration: a 60-second “sit still” step tunes an empty-vs-occupied SNR reference for the room and position. (This is setup calibration — not per-home model retraining; no LoRA-style adaptation is used.)

- Per-subcarrier amplitude extraction; configurable window/stride.

| ## 3. Breathing-Rate Estimation (On-Device DSP) | CORE |
| --- | --- |

Replaces Wi-Netra 6.3's ML classifier — a measurement task, so no trained model is used. The old “validity gate” is folded in here as the confidence output.

Description

Converts each CSI window into a breaths-per-minute value on the ESP32-S3 using signal processing kept to MCU-efficient operations (no on-device PCA). Rather than emitting a single fragile number, it emits BPM plus a confidence, and refuses to report a rate when the signal isn't trustworthy — outputting “no valid breathing” instead of a wrong value. This robustness is what separates a product from a lab script.

Key Features

- Hampel filter for impulse/spike removal.

- Subcarrier selection — concrete method: compute each subcarrier's signal variance within the 0.1–0.5 Hz band over the recent window and keep the top-K most responsive; optionally fuse selected subcarriers with CSCR (cross-subcarrier ratio) — the SA-WiSense technique that cancels CFO/SFO offsets and lets single-antenna hardware sense respiration at all.

- Phase decision committed: amplitude-only by default (single-antenna raw phase is corrupted by CFO/SFO/PBD, same reason the FYP is amplitude-only); phase is used only inside CSCR ratios where those offsets cancel — never raw.

- Butterworth band-pass to the respiration band (0.1–0.5 Hz, i.e. 6–30 breaths/min).

- Dual estimator — FFT + autocorrelation: report BPM as the median of the FFT peak and the autocorrelation peak, flagging disagreement as low confidence. Autocorrelation is robust to the non-sinusoidal, harmonic-rich shape of real breathing, preventing the classic error of locking onto the 2nd harmonic and reporting double the true rate.

- SNR gating & confidence output: windows below an SNR threshold return “no valid breathing” rather than a guess; the confidence value also disambiguates “empty room” from “possible apnea” before any alert fires.

- Sliding-window re-estimation for near-real-time output; runs fully on-device (any heavier fusion, e.g. PCA, is deferred to the phone).

| ## 4. Real-Time Cloud Sync | CORE |
| --- | --- |

Replaces Wi-Netra 6.6's MQTT layer — no local broker exists once the Pi is removed.

Description

Relays each device's computed breathing rate, confidence, and status from the ESP32-S3 to the Firebase Realtime Database over a persistent HTTPS connection; apps subscribe per device for live updates. Only low-rate derived values are sent — raw CSI never leaves the device.

Key Features

- Pinned cadence: breathing rate/status written every ~5 seconds; anomaly flags pushed immediately — roughly 12 writes/minute per device.

- Persistent connection: uses the async FirebaseClient library's long-lived connection mode, avoiding a costly TLS handshake on every write.

- The network push runs as a separate FreeRTOS task from CSI capture, so cloud I/O never introduces sampling jitter into the signal pipeline.

- Each device writes to its own Realtime Database path, enabling per-patient subscription and the multi-patient switcher.

- Security rules restrict writes/reads to authenticated devices and their linked accounts.

- Works from behind a home router with no static IP or port forwarding (outbound only).

| ## 5. Tiered Respiratory Anomaly Detection | CORE |
| --- | --- |

Tier 1 is Core and on-device; Tier 2 is a Stretch goal on the phone. Layered so the demo never depends on the hardest part.

Description

Detects abnormal breathing and surfaces it in the apps, escalating from deterministic rules to an optional learned model. The rule tier alone is a complete, defensible deliverable; the model is additive and honestly scoped as future work.

Key Features

- Tier 1 — Rule-based flags (CORE, on-device): Thresholds on the DSP breathing rate — apnea (no valid breathing for a set duration), tachypnea (above upper bound), bradypnea (below lower bound). Deterministic, explainable, no training data. Uses Module 3's confidence to avoid firing a false apnea alert on an empty room.

- Temporal voting (CORE): An anomaly must persist across consecutive windows before an alert is raised, suppressing spurious flags.

- Tier 2 — Unsupervised anomaly model (STRETCH, on phone): A small TensorFlow Lite autoencoder on the phone, trained on normal breathing, flagging high reconstruction error to catch waveform-shape anomalies (e.g. Cheyne-Stokes, apnea-like pauses) that rate thresholds miss. Honestly scoped: reaching it requires the rule tier fully working, a plan to collect labelled normal-breathing data from this setup, and real shape-anomaly test cases (not just breath-holding, which Tier 1 already catches). Positioned as “designed for future work, prototyped if time permits.”

- Scope note: A research/monitoring prototype — not a certified diagnostic or medical device.

| ## 6. Mobile App — Multi-Patient Monitor | SUPPORTING |
| --- | --- |

Single App User experience; whose data is shown is chosen by the patient switcher.

Description

A viewer app presenting live breathing rate and analytics, styled like a real-time patient monitor. One account can watch several patients — each on its own device — and switch between them, so a caregiver monitoring two family members sees each as a distinct live feed.

Key Features

- Patient switcher: select among the devices/patients linked to the account (e.g. Patient 1 → device A, Patient 2 → device B).

- Live oximeter-style readout: current breaths-per-minute with a real-time breathing trend and a clear “no valid breathing / signal low” state driven by Module 3's confidence.

- Historical analytics per patient: session logs, nightly/daily trends, rate distribution.

- Anomaly event timeline with severity and timestamp.

- Hosts the optional Tier 2 TensorFlow Lite model for on-phone shape-anomaly detection.

- Device pairing and per-device connection/health status; subscribes to each device's Realtime Database path.

- Submit a complaint / support request: App Users can raise a complaint from within the app (e.g. device or service issue), tagged with the account/patient; complaints sync to Firebase for the Admin panel to review and resolve.

| ## 7. Notification & Alert Management | SUPPORTING |
| --- | --- |

Adapted from Wi-Netra 6.7.

Description

Pushes alerts to the App User when a respiratory anomaly is confirmed, identifying which patient triggered it.

Key Features

- Firebase Cloud Messaging (FCM) push on a confirmed anomaly (after temporal voting), tagged with the patient/device.

- Severity levels (informational vs. urgent, e.g. suspected apnea).

- Alert history with acknowledgement/dismissal, per patient.

- Away-from-home delivery via Firebase as the bridge.

| ## 8. Admin Panel (Oversight) | SUPPORTING |
| --- | --- |

Deliberately small — a monitoring/oversight console, not the full FYP operations backbone.

Description

A small web panel (React) that makes the system feel like a real service: an operator can see every device’s health, manage users and device–patient assignments, review alerts across all patients, and handle complaints raised by App Users. Organised as five pages — Statistics/Analytics (landing), User Management, Alerts, Complaints, and Settings. Business-operations machinery beyond these is intentionally out of scope for the prototype.

Key Features

- Statistics / Analytics (landing page): fleet-wide breathing status across all monitored patients, live online/offline device counts, and aggregate anomaly trends.

- Device fleet view: each device’s online/offline and health status.

- User Management page: manage users and assignment of devices↔patients↔App User accounts (the links that power the app’s patient switcher).

- Alerts page: alert / anomaly oversight log across the fleet, with per-patient filtering and acknowledgement.

- Complaints page: review and resolve complaints submitted by App Users from the mobile app — view, assign status (open / in-progress / resolved), and respond. Complaints are stored in Firebase.

- Settings page: admin-level configuration — alert thresholds/defaults, account and role settings, and panel preferences.